New Theoretical Concept for Accurate Calibrations of Capacitance Boxes

M. Helmy A. Raouf

Department of Electrical Quantities Metrology

National Institute of Standards (NIS), Egypt

Abstract

Capacitance boxes are widely used as reference standards for calibration or construction of some measuring devices. This paper introduces new theoretical concept which represents a highly accurate method for calibration of decade capacitance boxes. It depends on nearly canceling or eliminating the systematic errors from the calibration process as clearly described. The values of all capacitance steps of a decade box are obtained in terms of one calibrated standard capacitance. All equations and steps of this new accurate concept are derived and demonstrated in details through this research.

1. Introduction

Any capacitance box consists of some decades which are usually formed by ten elements as illustrated in Fig. 1.

C1C2 C3C4C5                           C6C7C8 C9C10

Selecting

switch

moving

arm

Output terminals

Fig. 1: Block diagram of decade capacitance.

These boxes are used extensively as standards for capacitance measurements and as main components in some measurement devices of capacitance [1].The most commonly used bridges which use these decade boxes are, admittance ratio bridge (parallel resistance capacitance bridge) [2], Schering bridge [3], and others [4].  Series of decade capacitors are used with some bridges to determine the value of the 10 pF standard capacitance and higher capacitances values [5]. Also, calibrations of dissipation factor standards are mainly done by decade boxes [6].

The usual method used for calibrating a decade capacitance box is based on measuring its steps on a calibrated device [3]. Then, a correction factor is added to each step value according to the used range in this device. This method is not the most accurate method, because this correction factor is used for the whole device range rather than deduced each step value. In addition, this correction factor may be varied with time and measuring conditions, such as, temperature, humidity, connection wires and leads position…etc.  In addition to the lower accuracy, this method takes much effort and time. This is due to evaluation of all sources of errors and taking their values into account when the actual values of the measured decades steps are computed.

The new theory of decade capacitance boxes calibrations that introduced in this paper represents the most accurate and precise method of calibration. Hence, it depends on nearly eliminating or canceling the systematic errors from the calibration process and gives all values of the capacitance steps in terms of one calibrated standard capacitance. Therefore, the capacitance values of the decade box steps are investigated with the least errors and consequently minimum uncertainties [7].

2.  Theory   for   Accurate   Calibrations   of  Capacitance

Boxes

At first, supposing that, capacitance box which has a front panel shown in Fig. 2, has four decades and each decade has ten steps. Decade (I) has the lowest value of capacitance per step. Decade (IV) has the highest value of capacitance per step.

## I    		    II		III		IV

Fig. 2:  Front panel of capacitance box.

In order to calibrate this capacitance box by applying this new theory, we must have a standard capacitance having a nominal value, which equals to the nominal value of the first step of the lowest decade value (I). In addition, the actual value of this standard capacitance must be known accurately through the calibration at a higher precise standard laboratory.       Then, the calibration process of this capacitance box is started from the first step of the lowest value decade.

2.1   Definitions of the used  Symbols

Cs– measured value of the standard capacitance, while the actual value of this standard capacitance can be expressed as[8]:

Cs'=Csn+ds,    							                     (1)

where :

Cs'− actual value of the standard capacitance,

Csn− nominal value of the standard capacitance,

ds− deviation in the value of the standard capacitance from its nominal

value.

CIn  - nominal value of each step in decade I ,

CIIn - nominal value of each step in decade II ,

CIIIn- nominal value of each step in decade III ,

CIVn - nominal value of each step in decade IV .

CI1,CI2,....,CI10   -measured values of the steps in decade I, CII1,CII2,.....,CII10      - measured values of the steps in decade II,

CIII1,CIII2,......,CIII10  - measured values of the steps in decade III,

CIV1,CIV2,......,CIV10   - measured values of the steps in decade IV.

CI1',CI2',......,CI10'        - actual values of the steps in decade I,

CII1',CII2',.....,CII10'      - actual values of the steps in decade II,

CIII1',CIII2',......,CIII10'  - actual values of the steps in decade III,

CIV1',CIV2',......,CIV10'   - actual values of the steps in decade IV.

2.2  Calibration Steps

- The standard capacitance is connected to a suitable measurement device and its value  is measured, then we have [8]:

Cs=Csn+ds=Csn(1+dsCsn),                                                      (2a)

or,

Cs =Csn(1+Δs),        	   (2b)   						(5-2b)

where,

Δs=dsCsn =Cs−CsnCsn                                                                (2c)

which is a dimensionless number.

- Connect the decade capacitance to the measurement device and measure the value of its first step in decade (I) with the same measuring conditions of step 1, as: device settings, sensitivity, humidity, temperature,……., therefore :

CI1=CIn+dI1=CIn(1+dI1CIn)	,	      			         (3a)

or,

CI1=CIn(1+ΔI1)                                                                       (3b)

where

ΔI1=dI1CIn= CI1−CInCIn                                                               (3c)

consequently,

dI1s=CI1−Cs=CIn(1+ΔI1)−Csn(1+Δs),                              (4)

since,CIn=Csn, then,

CI1−Cs=Csn(ΔI1−Δs).                                                           (5)

Therefore, (ΔI1−Δs) represents the deviation between the capacitance value of step I1, and the capacitance value of Cs without any errors coming from measuring conditions or measurement device errors.

- The actual capacitance value of the step I1 , is computed in terms of the actual value of  standard capacitance Cs' , as:

CI1'=Cs'+dI1s,							        (6a)            or,

CI1'=Cs'(1+ΔI1s),						        (6b)

where,

ΔI1s=dI1sCs'. 						  	        (6c)

Then, the actual capacitance value that named Cs2' and equals to (Cs'+CI1') can be considered as a standard capacitance value to calibrate the second decade step I2. Hence, the deviation in its actual value from its nominal value   can be determined as illustrated in the following step.

- Connect the standard capacitances (Cs+CI1) in parallel to the measurement device and measure the equivalent capacitance value of their summation then:

Cs2=Cs+CI1=Csn2+ds2 ,      			                   (7a)

or,

Cs2=Csn2(1+Δs2), 						        (7b)

where,

Δs2=ds2Csn2.                                                           		        (7c)

- Connect the decade capacitance to the measurement device and measure the value of the second step in decade I, therefore:

CI2=CIn+dI2,				  	 	                   (8a)

or,

CI2=CIn(1+ΔI2), 					                  (8b)

where,

ΔI2=dI2CIn. 							                   (8c)

Accordingly,

dI2S2=CI2−Cs2, 				                             (9a)

then,

dI2s2=CIn(1+ΔI2)−Csn2(1+Δs2), 			        (9b)

since,CIn=Csn2,then:

dI2s2=Csn2(ΔI2−Δs2).					        (9c)

Therefore (ΔI2−Δs2) represents only the deviation between the capacitance value of step I2 and the capacitance value of CS2.

- The actual capacitance value of step I2 is computed in terms of the actual value of the standard capacitance value,Cs2'as:

CI2'=Cs2'+dI2s2, 						       (10a)

or,

CI2'=Cs2'(1+ΔI2s2)					                 (10b)

where,

ΔI2s2=dI2s2Cs2' . 						                 (10c)

Then, the standard actual value that named Cs3' and equals to (Cs'+CI2') can be considered as a standard capacitance value to determine the deviation in the third decade step I3 as the same calibration procedure of the step I2, and so on.

- Finally, CI10' is computed in terms of the actual standard capacitance value that named Cs10' and equals to (Cs'+CI9') as,

CI10'=Cs10'(1+ΔI10s10). 					                   (11)

- The last capacitance step of the first decade that named  CI10 is used as a standard capacitance to calibrate the first step in the second decade II that named CII1 ,hence they are nominally equal. Then the actual capacitance value CI10' is also, used to compute the actual value that namedCII1', of the step CII1. Also, the capacitances (CII1+CI10) are used in parallel to calibrate CII2, and so on …etc.

Consequently, by using this new calibration theory, all of the decades steps, which form the capacitance box, are calibrated in terms of the accurately calibrated standard capacitance.

3. Conclusions and Discussion

Capacitance decade boxes are used as reference standards for capacitance measurements and as main components in many measurement devices of capacitance. A highly accurate method for calibrations of decade capacitance boxes is introduced in this paper. It depends on nearly canceling or eliminating the systematic errors from the calibration process. The values of all capacitance steps of the decade box are obtained in terms of one calibrated standard capacitance. According to this new accurate theory, the capacitance values of the decade box steps are obtained with minimum errors and uncertainties which will be realized experimentally in future researches.

4. References

- Ibrahim M. H. Saad, A. B. Kotb, F. M. El-Lithy, and M. Helmy A. Raouf , “A  New  Proposed  Design  for  Capacitance Measuring  Bridge”, Al-Azhar University Engineering Journal, AUEJ, Vol. 2, No. 5, PP. 39-47, April 2007.

- Fluke Corporation, Legal Department, Everett, WA 98206-9090,

“Calibration: Philosophy in Practice, Second Edition”, May 1994.

- A. K. Sawhney, and Puneet Sawhney, “A Course in  Electrical and  Electronic Measurements and Instrumentation”, Dhanpat Rai   & Co.(P) LTD., 2002.

- H.Eren, “ Capacitance and Capacitance Measurements ” , CRC

Press LLC, Copyright 2000, http://www.engentbase.com.

- A. D. Inglis, B. M. Wood,  M. Côté, R. B. Young, and M. D. Early, “Direct Determination of Capacitance Standards Using a Quadrature Bridge and a Pair of Quantized Hall Resistors”, IEEE Trans.  Instrum.  Meas., Vol. 52, No. 2, PP. 559 -562, April 2003.

6.   E. D. Simmon,  G. J. Patrick, and  O. Petersons,  “ Calibration   of

Dissipation   Factor   Standards”,   IEEE Trans.   Instrum.  Meas.,

Vol. 48, No. 2, PP. 450 - 452, April 1999.

- “Guide to the Expression of Uncertainty in Measurement”, Saudi                    Arabian Standards Organization (SASO), 2006.

- M. Helmy A. Raouf, “Precision Measurements of Reactive Components   and   Modification   of  Some  Measuring  Devices”,  Ph. D. Thesis, Faculty of Engineering, Al-Azhar University, Cairo, May 2007.